## Slide 1
Thomas MARION, Operation group, Synchrotron SOLEIL
SOLEIL Tools to Improve the Reliability of the Accelerators

## Slide 2: Layout
SOLEIL in a Nutshell
Issue reporting and knowledge database
Web based data-monitoring system
Example of a Recent Major Incident
Conclusions

## Slide 3: SOLEIL in a Few Metrics
Nominal energy: 2.75 GeV
354 m electron Storage Ring (SR)
First user beam delivered in 2007
29 beamlines
Maximum intensity 500 mA, 4 filling modes
User beamtime: 5000 hours per year
About 700 publications a year
In the TDR phase of the SOLEIL II upgrade
Synchrotron SOLEIL is the 3rd generation French synchrotron radiation light source located South of Paris.

## Slide 4: SOLEIL in a Few Metrics
Beam availability: 97.32 % in 2023
113h of failure including 72h for a single failure
Mean Time Between Failure (MTBF): 146 hours
Cumulative beam time without failure 640h (4.5 weeks)
19 weeks > 99% availability
10 weeks = 100%
Chart: Efficiency during beamlines and RP sessions in 2023 4130 hours of beamtime delivered represent a beam availability of 97,32 % with a MTBF (Meantime between failures) of 146h20
| 2007 | 2008 | 2009 | 2010 | 2011 | 2012 | 2013 | 2014 | 2015 | 2016 | 2017 | 2018 | 2019 | 2020 | 2021 | 2022 | 2023
Available beam | 0.93828 | 0.95697 | 0.96395 | 0.96271 | 0.98422 | 0.96112 | 0.98 | 0.985 | 0.9889 | 0.93241 | 0.98698 | 0.9885 | 0.9886 | 0.9879 | 0.9841 | 0.9895 | 0.97322
 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
MTBF | 0.90972 | 1.3333 | 1.4375 | 1.7153 | 1.9583 | 2.0625 | 2.8264 | 3.1319 | 4.3958 | 3.3333 | 3.8611 | 3.6667 | 3.9951 | 4.375 | 4.5417 | 5.8125 | 6.0972

## Slide 5: Historical Ecosystem
Operation Group
Operation LogBook
Group LogBook
Statistics
Procedure
Procedure
Intervention
Vacuum Group
Group LogBook
Procedure
Intervention
Disadvantage :
  Each group have their own ecosystem
  Data are scattered
  Slow interaction
Power Supply
Group

## Slide 6: Ecosystem
JIRA Incidents
Statistics
Confluence
JIRA Intervention
Grafana
JIRA: Incident Tracking Software
CONFLUENCE: Collaborative wiki system
GRAFANA: Data-monitoring web-based software
Goals:
  Make interaction between group easier
  Improve traceability and reporting at group level
  Save time
  Create a Knowledge Database
Accelerator DB
Operation Group
Vacuum Group
Power Supply
Group

## Slide 7: Issue Reporting and Knowledge Database

## Slide 8: Incident Reporting: From Elog to JIRA
Until 2018: Elog
  Manually by Operator
  Slow
  Hard to request with scripts
  No collaboration between group
  Incident resolution transmitted orally to the operator
    Can lead to information loss or transcription error
Since 2018: JIRA
  An issue-tracking tool developed by Atlassian
  At SOLEIL
    Since 2015 for IT Teams
    Since 2018 for Machine Incidents
  Objectives:
    Cross-team collaboration
    Share operation knowledge

## Slide 9: Incident Workflow
Known Incident not requiring the intervention of the support group (INIT, RESET, ...)
Creation
Project: ServiceDesk CTRLDESK)
type of request: Incident
components: Machine operation
Summary:
Description:
Operator Rooting
Resolved
Level 1
Closed
Incident requiring action of the support group
Incident requiring action of the Operator
Closing the ticket by the Incident Manager
Create / link to problem(s)
Create / link : Mondays Machine or shutdown intervention
Reviewed by support group
Add comments / report ... by the support group
Start date/ End Date
Beam incidence
Group (support) / equipment / fault
Location
Level 2

## Slide 10: JIRA Dashboards Tailored for Accelerator and IT Support Groups
Incidents in progress - To be processed by the technical team
Incidents in progress – with impacts on the team
Resolved issues assigned to the group - waiting for REVIEW by the team
Problems
Incidents closed (resolved and reviewed) over 365 rolling days
Resolved issues – with impacts on the team

## Slide 11: Weekly Operation Meeting
Incidents:
  We review all the incidents of the previous week
  Time saved for preparation of the meeting for the organizer
  Better efficiency during the meeting
    Support group can prepare in advance
Intervention: urgency, priority, duration
  We plan the next Monday's intervention shift
  and the technical shutdown intervention
One-page incident summary sorted by impact
Notes: JIRA is the support of the meeting

## Slide 12: JIRA User-Friendly and Efficient Interface for Report and Incident
Goals
  Report everything and Reducing the number of manual actions
  Ease the task to report as soon as possible an incident with all relevant information and a link to our operation Elog
  Redirect incidents and escalations to the right support group
Development
  Web interface
    Fast incident creation
    Avoid error in rooting
  Automatic issue creation (Low Current, Insertion device in fault…)
    No misses or errors
    Log attachments

## Slide 13: Web Statistics
  Developed in Python
  Build on top of Jira
  Updated every day
  Statistics since 2008
  Accessible to every group

## Slide 14: Intervention Tracking
Historically:
  Handwritten forms for downtime interventions.
  Mail for intervention on machine day
Disadvantage:
  Time consuming
  Difficult to track
  The applicant, security and radio-protection need to be physically in the control room to sign the paper
  Hard to archive
Integration in JIRA:
  Easier to track and archive
  Requests validation and signatures made easy
  Reliable Metric for annual review
  Link between incident and intervention

## Slide 15: Confluence
  Collaborative software based on a wiki system.
    Procedure
    Meeting report
    Decision Logbook
Operator reaction time improved

## Slide 16: JIRA and Related tools Benefits
Benefits
  Time-saving for reporting
    ≈ 700+ Tickets / year -> 50% Automatic
  Time-saving for Statistics and weekly meeting preparation
  More information for support group
    Significant reduction of minor incident
  Incident resolution time has decreased drastically
  Traceability has increased
Quality insurance
  Incident manager (Operation Engineers)
  Daily incident review to consolidate the information
Chart: Number of JIRA Ticket created since 2018
| 2018 | 2019 | 2020* | 2021 | 2022 | 2023*
Total incident + Beam | 1335 | 1139 | 791 | 756 | 714 | 591
Automatically created ticket | 694 | 532 | 382 | 384 | 352 | 288
Incident
Procedure
Intervention
Knowledge database

## Slide 17: Grafana
Web based data-monitoring system

## Slide 18: Grafana
What is Grafana
  Grafana is a free web-based software that allows you to build your own dashboards on various data sources.
  We developed a data source plugin dedicated to our archiving system.
Advantages
  “Live” data
  Accessible from a web browser on any network
  Every user have their own space
  Dashboard creation is fast, No-Code and easy to share
Grafana dashboard creation

## Slide 19: Examples of Grafana Dashboards

## Slide 20: Grafana Use Cases
On call monitoring:
  Before:
    Connect to the Machine Network
    Launch archiving plot / or dedicated application
Emergency development:
  After the inauguration of our new water-cooling station the utility group had a problem reading back values from their supervision software.
  We developed a dashboard in 2 hours on Friday evening, so that they could monitor the data remotely.
  With Grafana:
    In the same view we can monitor all the metrics we want on the same dashboard
    Visible from a Computer or a Smartphone
Data monitoring in and outside of control room

## Slide 21: Example of a Recent Major Incident
Transfer Line Bending Magnets incident

## Slide 22: Transfer Line BMs (Bending Magnets) incident
About 20 minutes later: Fire Alarm in the Linac area. Smoke is coming out of the Bending Magnets.
We turned the Power supply OFF. Water cooling is turned ON again.
cooling tubes have burst due to the temperature drop.
The BM coils seems to be severely damaged
The Power supply was not turned OFF by the security box,
Repair intervention on a Booster RF cavity on a holiday (reduced personnel). Not possible to locally turn OFF the water cooling, The entire Booster cooling circuit is turned OFF.

## Slide 23: Incident Timeline
Incident
Monday (public holiday)
Tuesday
Wednesday
Thursday
Friday
Emergency meeting with On-Call expert:
Decision to cancel user beam on Tuesday morning
Emergency meeting involving groups expert
Control and interventions
Interlock system
Hydraulic system
Power supply
Magnets
Vacuum
Accelerator restart + live testing
Beam delivery
Tuesday meeting:
We only have two spare coils (and we need 4)
How long to change them? (1 week)
Are the coils still viable?
Why did the interlock system not work?
Could the problem exist elsewhere?
Wednesday evening meeting
Decided to restart the Machine on the next day to deliver the beam on Friday morning.
.
Yes
1 more defective security box
1 relay of power supply in test mode
Damaged coils replaced during summer shutdown
Periodic test of all the safety box at every shutdown
Lessons learned for Magnet Protection System upgrade (With PLC)

## Slide 24: Incident Conclusions
  Complex incident involving 5 groups
  All actions have been tracked in JIRA and Confluence
    Incident
    Tests and Interventions report
    Meeting notes
  Smooth information flow between groups
  72 hours of beam unavailability
    But it could have been much worse

## Slide 25: Conclusions and Perspectives

## Slide 26: Conclusions and Perspectives
Enhanced Operation
  How could Machine Learning improve operation reliability?
    Challenges of SOLEIL II with a significant increase of equipment and more complex accelerators handled, serviced, and operated with the same human resources
    Predictive Maintenance
    Benefit of Retrieval-Augmented Generation (future of LLM and generative IA) using our high-trusted knowledge databases
Conclusion:
  Better interaction between groups
  Traceability and Efficiency increased
  Incident resolution time has decreased
  Creation of a high-trusted knowledge Database
Chart: Efficiency and MTBF during beamlines and RP sessions since 2018
| 2018 | 2019 | 2020 | 2021 | 2022 | 2023
Faisceaux disponibles | 0.9885 | 0.9886 | 0.9879 | 0.9841 | 0.9895 | 0.97322
 | 0 | 0 | 0 | 0 | 0
 | 0 | 0 | 0 | 0 | 0
MTBF | 3.6667 | 3.9951 | 4.375 | 4.5417 | 5.8125 | 6.0972

## Slide 27: Thank you for your attention
Questions?