## Slide 1: Multimodal DAQ system at Balder beamline
Heps-max iv visit 2025
Software group
Vanessa Silva
vanessa.silva@maxiv.lu.se

## Slide 2
Software group
Balder
Hard X-ray Spectroscopy Beamline

## Slide 3: Balder
Hard X-ray Spectroscopy Beamline
Software group
Notes: The Balder beamline is dedicated to X-ray absorption and emission spectroscopy in the medium to hard X-ray
energy range (2.4 to 40 keV) and is located at the 3 GeV ring. This is the only spectroscopy beamline in the
hard X-ray energy range at MAX IV, while several soft X-ray spectroscopy beamlines exist.

## Slide 4
Main experiment performed at Balder
Combined
XAS
XRD
X-ray Absorption Spectroscopy
X-ray Diffraction

## Slide 5
Main Application and Project Scope
Combined
XAS
XRD
Hard X-ray Beamline
No need for in vacuo sample environment
Easy to perform in situ/in operando experiments
Perform XAS (range 1000 eV) in sub-second time resolution
Quickly change edges several thousands of eV apart: suitable for multi-element materials
Perform XRD combined with XAS in the end of every edge scanning
System flexible enough to allow user to configure desired experiment combination order
Data saved on the same HDF5 file
The experiment scan and its multiple energy edges & techniques are configured only once by the user in a single scan macro
Battery Sample Environment @ Balder
Multimodal Experiment

## Slide 6: The challenge
Scan motion orchestration done by hardware
Allow multi-techniques scan to be performed
Easy to be configured by the user
Use Sardana framework
Software group
New concept of scan macro
Current
Scan macro
(continuous scan)
Scan single energy range
Prepares and acquire data using same pulse train for all detectors
New
Scan macro
(continuous scan)
Scan multiple energy ranges
Prepares and acquire data using different pulse trains for different detectors and its associated technique
Notes: Scan multiple energy ranges
Prepares and acquire data using different pulse trains for different detectors and its associated technique

## Slide 7: How’s the beamline energy scanned?
Software group

## Slide 8
Energy Scan
Balder’s Optical Hutch

## Slide 9
Energy Scan
Balder’s Optical Hutch

## Slide 10
mono beam
white beam
Si (111)
vertical
beam offset
Bragg motor
DCM
Balder’s Optical Hutch
Energy Scan

## Slide 11: ç
Si (111)
white beam
mono beam
vertical
beam offset
Bragg motor
DCM
What’s the correlation between Bragg Angle θ and energy?
Bragg’s Law
Energy Scan

## Slide 12: ç
Motion Control
Si (111)
white beam
mono beam
vertical
beam offset
Bragg motor
DCM
performs parameterized trajectories
EtherCAT master motion controller
ACS SPiiPlusEC
master-follower implementation by synchronization formula for fixed-exit
allow implementation arbitrary motion trajectories for energy scans
Full energy edge can be executed up to a 2 Hz repetition frequency
Monochromator Motion Control
Notes: The two axes, which are used for energy scanning (Bragg angle and perpendicular motion of the second crystal
- x2perp) are driven by a high-precision and high-dynamic range dedicated motion control system (Figure 10).
Motions are controlled by an EtherCAT master motion controller (ACS SPiiPlusEC) with a motion control
update frequency of 5kHz

## Slide 13: ç
Motion Control
Si (111)
white beam
mono beam
vertical
beam offset
Bragg motor
DCM
EtherCAT
Load different types of trajectory into the controller buffer, based on pre-defined and programmed functions
Monochromator Motion Control Hardware Integration
Notes: The two axes, which are used for energy scanning (Bragg angle and perpendicular motion of the second crystal
- x2perp) are driven by a high-precision and high-dynamic range dedicated motion control system (Figure 10).
Motions are controlled by an EtherCAT master motion controller (ACS SPiiPlusEC) with a motion control
update frequency of 5kHz

## Slide 14: How’s the experiment synchronization performed?
Software group

## Slide 15
How’s the data acquisition performed?
Si (111)
white beam
mono beam
vertical
beam offset
Bragg motor
DCM
performs parameterized trajectories
ACS SPiiPlusEC
motor controller
TTLOUT
encoder reading input
Analog input
PandABox
Detectors
Eiger
Alba EM
Trajectory
Synchronization
DAQ

## Slide 16
Position-based hardware synchronization
PandABox
TTLOUT
encoder reading input
Analog input
Bragg Motor

## Slide 17
Position-based hardware synchronization
PandABox
Standard latency time = 0.1 ms
Why gates?

## Slide 18
Main experiment performed at Balder
Combined
XAS
XRD
Coming back to the beamline techniques ….

## Slide 19
Main experiment performed at Balder
Combined
XAS
XRD
Coming back to the beamline techniques ….

## Slide 20: z
Main experiment performed at Balder
XAS
DCM motion trajectories with EtherCAT ACS
ALBA Electrometer

## Slide 21
Main experiment performed at Balder
Combined
XAS
XRD
Coming back to the beamline techniques ….

## Slide 22
Main experiment performed at Balder
XRD
Fixed energy positions
beginning and/or end of XAS spectrum
Eiger

## Slide 23
Position-based hardware synchronization
TTLOUT
encoder reading input
Analog input
Bragg Motor
Generate a gate for XRD fixed energy position [beginning or end of the spectrum]
Generate gates for all configured energy positions for XAS intensity measurements & PCAP
XRD
XAS

## Slide 24
Position-based hardware synchronization
TTLOUT
encoder reading input
Analog input
Bragg Motor
Bragg Position
PCAP
Alba EM
Eiger
Time

## Slide 25: How’s the experiment configured and controlled at the beamline?
Software group

## Slide 26
Pool
MacroServer
Motors
1D Channels
2D Channels
Counters
Controllers
Macros
Recorders
Experiment Control with Sardana

## Slide 27
Pool
MacroServer
SPOCK
🖖
Experiment Control with Sardana

## Slide 28
START DATA RECORD
PREPARE EQUIPMENTS
MOTION ORCHESTRATION
RUN PRE-SCAN HOOKS
END DATA RECORD
RUN POST-SCAN HOOKS
SYNCHRONIZATION
peristaltic pump
detector
DATA ACQUISITION
Experiment Control with Sardana
Scan Orchestration
User input layear
scan macro

## Slide 29
MOTION ORCHESTRATION
ç
Si (111)
white beam
mono beam
vertical
beam offset
Bragg motor
DCM
EtherCAT
Load different types of trajectory into the controller buffer, based on pre-defined and programmed functions
Experiment Control with Sardana

## Slide 30
START DATA RECORD
PREPARE EQUIPMENTS
DATA ACQUISITION
RUN PRE-SCAN HOOKS
END DATA RECORD
RUN POST-SCAN HOOKS
SYNCHRONIZATION
peristaltic pump
detector
ACS
Experiment Control with Sardana
Scan Orchestration
User input layear
scan macro
Run buffer program -> start trajectory
Notes: start

## Slide 31
hmoscan <config_file> <nb_repeats> <latency_time:Optional>
JSON format
JSON format
Different technique have different parameters defined by the keys
Complex scan parameters can be defined easily in a JSON file format
Experiment Control with Sardana
hardware motion orchestrated scan

## Slide 32: Configuration file in JSON format example
It can be edited by beamline staff and/or users
Different techniques can be sequenced (XAS, XRD and EXAFS)
Software group
Experiment Control with Sardana
hmoscan <config_file> <nb_repeats> <latency_time:Optional>
JSON format

## Slide 33
hmoscan <config_file> <nb_repeats> <latency_time:Optional>
JSON format
Taurus GUI where the USER can organize the sequence of techniques to be performed using “drag and drop” system.
Each technique will have its own set of parameters to be configured. The USER can choose between XAS scan and XRD. Inside each technique, many variations are allowed.
The GUI generates the experiment JSON file with specific keys
Experiment Control with Sardana

## Slide 34
hmoscan <config_file> <nb_repeats> <latency_time:Optional>
JSON format
BalderScans
PositionArray
TriggerArray
ConfigureACS
TotalTriggers
SPMTriggers
Experimental channels
detector
Trigger/Gate controller
ACS Tango Device
Load trajectory using pre-configured functions into ctrl BUFFER
Experiment Control with Sardana

## Slide 35
START DATA RECORD
PREPARE EQUIPMENTS
DATA ACQUISITION
RUN PRE-SCAN HOOKS
END DATA RECORD
RUN POST-SCAN HOOKS
SYNCHRONIZATION
peristaltic pump
detector
ACS
hmoscan <config_file> <nb_repeats> <latency_time:Optional>
BalderScans
PositionArray
TriggerArray
ConfigureACS
TotalTriggers
SPMTriggers
Experiment Control with Sardana
Run buffer program -> start trajectory

## Slide 36
Balder Use Case Tested with This System

## Slide 37: Initial Results
Software group

## Slide 38
XAS for Standard Cu Foil
Not HW orchestrated scan – 84 seconds
HW orchestrated scan
Up to 10x faster
Very good data quality as evaluated on last beamline review in 2024

## Slide 39
Main Application and Project Scope
Combined
XAS
XRD
Hard X-ray Beamline
No need for in vacuo sample environment
Easy to perform in situ/in operando experiments
Perform XAS (range 1000 eV) in sub-second time resolution
Quickly change edges several thousands of eV apart: suitable for multi-element materials
Perform XRD combined with XAS in the end of every edge scanning
System flexible enough to allow user to configure desired experiment combination order
Data saved on the same HDF5 file
The experiment scan and its multiple energy edges & techniques are configured only once by the user in a single scan macro
Battery Sample Environment @ Balder

## Slide 40: Next steps
Software group

## Slide 41
Full replacement of the Configuration Layer by Sardana new features
MultiSynchDescription: allow detectors to be configured for different pulse trains
Decouple data acquisition from data recording
Perform XAS-2d mapping with HW orchestrated scan
Next steps

## Slide 42: Thank you!
Software group
Vanessa Silva
vanessa.silva@maxiv.lu.se