## Slide 1: Emittance Diagnostics
NSLS II-MAX IV workshop
5-6 Oct. 2026
Jonas Breunlin on behalf of the MAX IV diagnostic beamlines team

## Slide 2: Outline
Intro : the MAX IV bending magnet diagnostic beamlines
Transverse and longitudinal beam diagnostics today
Towards MAX4U: resolving smaller horizontal beam sizes
Optics components degradation

## Slide 3: Overview: bending magnet diagnostic beamlines at MAX IV
Bend magn diagBLs:
B320B & B302B
Two beam lines, because we search two unknowns, Horizontal emittance and Relative energy spread , from two measured horizontal beam sizes .
Vertical emittance simply from:

## Slide 4: Longitudinal (time-resolved) beam diagnostics
Streak Camera
Time-Correlated Single Photon Counting
Importance of long. diagnostics bunch lengthening with multiple passive / active harmonic cavities, also for the upgrade!

## Slide 5: Transverse beam size diagnostics today
Imaging with synchrotron radiation from micrometer electron beams: Interference and SR inherent effects dominate!
Model-based calculations done in the Synchrotron Radiation Workshop (SRW)
Fringe visibility gives beam sizes
Online image evaluation for beam size measurements
Combination of beam size measurements to get emittances and energy spread:
Vertical intensity profile, sensitive to
B320B measurement
SRW calculation for B320B
Horizontal intensity profile, sensitive to

## Slide 6: Preparing for the upgrade
Source points of diagnostic beamlines are kept
Vacuum chambers will be designed if needed, such that vis-UV dipole radiation can be extracted.
Majority of the vacuum system etc. can be kept
Reduction in horizontal emittance, beta functions and dispersion.
  Resolve horizontal beam sizes > 10 micrometers
  Requires retuning of the horizontal beam size method
No significant changes to the vertical beam size and emittance expected.
Lattice functions at B302B source point for the 65 pm TDR-1 lattice. Åke Andersson
Example shown to the right:
SRW calculation of a 10 micrometer horizontal beam, imaged at 365 nm with a horizontal acceptance angle of 16 mrad

## Slide 7: Horizontal beam size methods
Beam size resolution from the slope of the sensitivity curve:
  change of the observable VPR (valley-to-peak ratio) per change in beam size
A horizontal method combines multiple beamline parameters.
The goal is to maximize beam size sensitivity for a given (range of) beam size(s).
Many beamline parameters can be easily changed:
  Wavelength -> bandpass filter
  Acceptance angle -> movable apertures
  …etc.
There is a limit to the maximum accepance angles! -> engineering challenge!
B320B today, optimized for 22 micrometers
B320B @ MAX4U, optimized for 10 micrometers. Two method candidates.
Side note: even today‘s method has a non-zero slope at MAX4U beam sizes!

## Slide 8: Horizontal beam size resolution estimates
The slope:
Δ(vpr) / Δ(σ_x/σ_x,design) = 0.27
std(vpr) / 0.27 = std(σ_x/σ_x,design) = 0.007
Comparing methods by their horizontal beam size resolution
Measured mean and standard deviation of B320B horizontal valley-to-peak ratio and σ_x at cold beam:
vpr = 0.1546 ± 0.0019 -> 1.3 %
σ_x = 22.934 ± 0.167 µm -> 0.7 %
The slope:
Δ(vpr) / Δ(σ_x/σ_x,design) = 0.175
std(vpr) / 0.175 = std(σ_x/σ_x,design) = 0.0108
The slope:
Δ(vpr) / Δ(σ_x/σ_x,design) = 0.151
std(vpr) / 0.175 = std(σ_x/σ_x,design) = 0.0126
-> The standard deviation of the measured horizontal beam size is expected to increase from 0.7% (B320B, R3) to 1.26% (B320B, MAX4U) assuming a more conservative method.

## Slide 9: UV-induced optics degradation
Fused Silica Lens (76 mm diameter)
After ca. 1 year of exposure. The visible contamination is Hydrocarbons
The optics degradation affects beam size measurements already after a short while of exposure: 2 weeks of delivery beam (400 mA).
Image by Eleni Moschovaki
Apart from surface contamination there is a bulk effect as well. Image obtained via HeNe illumination on a test bed on an B320B-exposed lense, that does not show any degradation visible to the eye!
2 weeks of delivery beam (400 mA).
Approach currently under investigation: protecting the optics by introduction of an additional mirror (downstream the SiC, but upstream of any FS element) and thereby removing harmful spectral components in the UV.

## Slide 10: Conclusions
vis-UV bending magnet diagnostic beamlines continue to be vital diagnostics…
worth taking them with us through the upgrade
there a challenges ahead:
  Engineering: preserve opening angles
  Opitimize the ‘methods’
  Understand and solve the optics degradation problem

## Slide 11: Complementary Material